## Slide 1: An Introduction to Software Failure Modes Effects Analysis (SFMEA)
Ann Marie Neufelder
© SoftRel, LLC, 2015
amneufelder@softrel.com
www.softrel.com
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Welcome to the online edition of Software Failure Modes Effects Analysis course by Ann Marie Neufelder of Softrel, LLC

## Slide 2: Agenda
Overview
Prepare the SFMEA
  Identify scope of SFMEA
  Identify resources required
Analyze the software failure modes and root causes
  Research past failure modes
  Brainstorm failure modes
  Analyze failure modes and root causes
Analyze consequences
  Identify subsystem, system effects, patient and user effects
  Identify preventive measures
  Identify severity and likelihood
Identify mitigation
  Identify corrective actions
  Identify compensating provisions
  Revise RPN
Avoid common mistakes
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: First we will cover a few basic things that you need to know to understand how to perform a software FMEA. Then the class agenda will follow the 4 basic steps of the software FMEA. The class will finish by illustrating a few common mistakes with regards to performing software FMEAs.

## Slide 3: Software is increasing in size
The increase in size of F16A to F35 is just one example[1]
With increased size comes increased complexity and increased failures due to software as shown next
1.0 Introduction
Chart: Size in SLOC of Fighter Aircraft since 1974
| 1974 | 1984 | 1997 | 2006 | 2012
 | 135000 | 236000 | 1700000 | 6800000 | 24000000
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Before we begin the software FMEA presentation, let’s start with explaining why the software FMEA has become one of the most popular software analyses. Over the decades, software has grown exponentially in size as shown in this figure. The size of an average software system makes it very difficult to test thoroughly and completely. Even medium sized software systems have an almost infinite number of possibilities with regards to test paths. Additionally many software failures are related to what the software does NOT do and SHOULD do. These are things that are often not in the test plan because they are not in the software requirements or design documents. The need for software FMEAs only increases as the size and complexity of the software system increases.

## Slide 4: Just a few failure events due to software
Failure Event | Associated software fault
Several patients suffered radiation overdose from the Therac 25 equipment in the mid-1980s. [THERAC] | A race condition combined with ambiguous error messages and missing hardware overrides.
AT&T long distance service was down for 9 hours in January 1991. [AT&T] | An improperly placed “break” statement was introduced into the code while making another change.
Ariane 5 Explosion in 1996. [ARIAN5] | An unhandled mismatch between 64 bit and 16 bit format.
NASA Mars Climate Orbiter crash in 1999.[MARS] | Metric/English unit mismatch. Mars Climate Orbiter was written to take thrust instructions using the metric unit Newton (N), while the software on the ground that generated those instructions used the Imperial measure pound-force (lbf).
28 cancer patients were over-radiated in Panama City in 2000. [PANAMA] | The software was reconfigured in a manner that had not been tested by the manufacturer.
On October 8th, 2005, The European Space Agency's CryoSat-1 satellite was lost shortly after launching. [CRYOSAT] | Flight Control System code was missing a required command from the on-board flight control system to the main engine.
A rail car fire in a major underground metro system in April 2007. [RAILCAR] | Missing error detection and recovery by the software.
1.0 Introduction
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Over the last 5 decades there have been many system failures due to software. This page shows just a few of them. Your book describes several of them. However, your book and this presentation only scratches the surface. For every software related event that is in the public domain it’s suspected that several more or not in the public domain due to security and confidentiality.

## Slide 5: Terms
Failure mode – The means by which the software could contribute to a system failure
Effect – Behavior that results from that failure
SFMECA – Software Failure Modes Effects Criticality Analysis includes a severity and likelihood ranking used for prioritizing corrective actions
Error - A discrepancy between a computed, observed, or measured value or condition and the true, specified or theoretically correct value or condition.
Defect – The manifestation of an error in the software product such as the requirements, design, or code. This is also a root cause for a failure event.
Fault – A defect that has resulted in one or more failures.
1.1 Software FMEA defined
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Before we start you will need to be familiar with a few of the terms used in this course.

## Slide 6: The viewpoint of a SFMEA
SFMEA
works
this way
Customer and software requirements, architectural design, interface design, code, users manuals
Failure modes and root causes applicable to incorrect requirements, design, code, users manuals
Immediate Effect (crash, hang, etc.)
Effect on subsystem (loss of data, communications between systems, etc.)
Effect on system (loss of system, degradation of system, downtime, etc.)
Effect on end users
Visible to SW engineers
Visible to users
Possibly visible to both
1.1 Software FMEA defined
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 7: Key benefits of Software FMEAs
SFMEAs are useful for identifying
  Defects that easier to see when looking at the design or code but difficult to see during testing
    i.e. can be used to improve the efficiency of design or code reviews
  Single point failures due to software
  Defects that cannot be addressed by redundancy or other hardware controls
  Abnormal behavior that might be missing from the requirements or design specifications
  Unwritten assumptions
  Features that need fault handling design
Addressing one failure mode could mean eliminating several failures
1.2 SFMEA Purpose
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Simply stated, people often overestimate how many and the types of defects that they can find during software and systems testing. The purpose of the software FMEA is to identify what the software should not do so that the requirements, design, code and test plans can reflect that. It’s normal for human beings to define requirements in positive terms. However, it is often the unexpected events that cause the software and hence the system to fail. This analysis provides for a way to identify the negative requirements that will ultimately require fault handling.

## Slide 8: SFMEA Limitations
Highly dependent on experience of analyst with system under development
Results depend on willingness of analyst to consider the failure space – what can go wrong
Dependent on adequate software requirements, design
Most productive when they focus on riskiest software with most likely failure modes
Not as effective when conducted too early or too late
1.4 SFMEA Limitations
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The SFMEA is a powerful analyses tool. However, it is dependent on the people who perform the analyses. If the analysts are willing and able to analyze what can go wrong with the software then the analysis can and will have a return on investment. Since software is developed by humans and all software defects are inherently caused by human mistakes made in the requirements, design and code, it can often be difficult for analysts to be objective when performing the analysis. In addition to the willingness and capability of the analysts, the software FMEA is also dependent on timeliness. It’s also very important that the analysis focus on the riskiest failure modes and the riskiest parts of the software to ensure that it’s effective. This class provides an entire section on how to plan the SFMEA to ensure that the above limitations don’t reduce the return on investment.

## Slide 9: Existing SFMEA guidance
Guidance | Comments
Mil-Std 1629A Procedures for Performing a Failure Mode, Effects and Criticality Analysis, November 24, 1980. | Defines how FMEAs are performed but it doesn’t discuss software components
MIL-HDBK-338B, Military Handbook: Electronic Reliability Design Handbook, October 1, 1998. | Adapted in 1988 to apply to software. However, the guidance provides only a few failure modes and a limited example. There is no discussion of the software related viewpoints.
“SAE ARP 5580 Recommended Failure Modes and Effects Analysis (FMEA) Practices for Non-Automobile Applications”, July, 2001, Society of Automotive Engineers. | Introduced the concepts of the various software viewpoints. Introduced a few failure modes but examples and guidance is limited.
“Effective Application of Software Failure Modes Effects Analysis”, November, 2014, AM Neufelder, produced for Quanterion, Inc. | Identifies hundreds of software specific failure modes and root causes, 8 possible viewpoints and dozens of real world examples.
1.5 Existing SFMEA Guidance
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The Military Standard on FMEA doesn’t discuss software at all. The military handbook discusses it but doesn’t provide the level of detail required to fully apply the FMEA to software. The SAE guidebook provides more detail but still shows very few software specific failure modes and guidance. This presentation is based on the latest guidebook published by Quanterion, Inc which is dedicated to providing the failure modes, viewpoints and examples needed for any organization developing or acquiring software systems to perform the software FMEA.

## Slide 10: Software Failure Mode Taxonomies
One major gap in previous general FMEA guidance is the lack of software failure mode taxonomies.
This class focuses on the failure modes and root causes that apply to virtually all software systems.
The SFMEA analyst is encouraged to brainstorm additional software taxonomies
  In 1983, Dr. Boris Beizer first published the “Taxonomy of software bugs” based on the development activity that introduces the “bug” [ST1].
  In 1989 Peter Neumann and Don Parker developed taxonomies for computer misuse [NEUMANN].
  In 1999 Dr. Robert Binder published a taxonomy on software faults relevant to object oriented development [ST2].
  In 2002 Giri Vijayaraghavan and Cem Kaner developed taxonomies for e-commerce related faults [ST3] [ST4].
  Mitre has published hundreds of common weaknesses which lead to security vulnerability [ST5]. Several of these weaknesses are also related to unreliable software.
  Other taxonomies have been developed for particular types of computers. [ST6][ST7]
  This author has compiled a listing of process and product related defects based on decades of writing and testing software and benchmarking development practices versus actual fielded defect density.[ST8] [ST9]
1.5 Existing SFMEA Guidance
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Prior to the publication of the “Effective Application of Software Failure Modes Effects Analysis” the available FMEA guidance did not provide sufficient guidance for the software failure mode taxonomies. This presentation provides software failure modes and root causes that apply to virtually all software systems. There are other taxonomies that apply to certain types of software systems. For example, there is a taxonomy for computer reuse, object oriented software, e-commerce software, software vulnerabilities, and specific types of computers. There is also a taxonomy written by this author concerning the hundreds of process related failure modes. The reader is encouraged to explore other taxonomies as applicable.

## Slide 11: The process for doing a Software Failure Modes Effects Analyses
1.6 Software FMEA Steps
Generate CIL
Mitigate
Analyze failure modes and root causes
Prepare the Software FMEA
Identify
resources
Brainstorm/
research
failure
modes
Identify
equivalent
failure modes
Identify
consequences
Identify local/
subsystem/
system
failure effects
Identify severity
and likelihood
Identify corrective
actions
Identify preventive
measures
Identify
compensating
provisions
Analyze
applicable
failure modes
Identify
root causes(s)
for each
failure mode
Generate a Critical
Items List (CIL)
Identify
applicability
Set
ground
rules
Select
viewpoints
Identify
riskiest
software
Gather
artifacts
Define
likelihood
and
severity
Select
template
and
tools
Revise RPN
Decide
selection
scheme
Define scope
Identify resources
Tailor the SFMEA
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The process for doing a software FMEA is similar to that for doing a hardware FMEA. The first step is to prepare the software FMEA. This step includes defining the scope of the software FMEA, identifying the resources needed for the software FMEA and tailoring the software FMEA to the particular needs of the project. The next major step is to analyze the failure modes and root causes. This is where most of the effort is typically spent. Once the applicable software specific failure modes and root causes are identified the consequences on the software and the system are identified for each failure mode and root cause. Then, the corrective action and mitigation for the failure modes is identified. The risk probability number is updated if the failure mode is mitigated or is planned to be mitigated. Finally the failure modes and root causes that are equivalent (if any) are consolidated and a list of Critical Items is generated. At this point the software and hardware critical items are typically merged so as to produce a system wide list of critical items. The CIL will often be used to enrich the existing test plans as well as the existing requirements and design documents. The CIL can also be used as inputs for any existing health monitoring software.

## Slide 12: Differences between hardware and software FMEA
Software failure modes are analyzed via different “viewpoints” than the hardware
Software has its own unique set of failure modes
A SFMEA can and will analyze how software reacts to a hardware failure
The same FMEA template can be used for software with some minor alterations
  There is a different template for each of the 8 software failure mode viewpoints
1.7 Differences between SFMEA and hardware FMEA
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: If you have performed a FMEA on hardware it’s useful to know the differences when applying it to software. Software is not going to fail due to wear out, temperature, vibration, etc. It will fail due to faulty requirements, faulty interfaces, faulty communications, faulty timing, faulty sequences, faulty logic, faulty data definition, faulty memory allocation, faulty installation, security vulnerabilities, etc. The software will have different viewpoints. A viewpoint is how you look at the software. There are failure modes that apply to any software system which will be unique from hardware failure modes. A software FMEA can analyze how the software reacts or should react to a hardware failure. However, keep in mind that the software FMEA doesn’t analyze the hardware failure, but rather how the software handles that failure. The similarities in the analyses are that the same template that’s used for a hardware FMEA can be used for software FMEA with a few minor alterations. On the next slide you will learn more about the viewpoints and failure modes…

## Slide 13: Outline for this section
  2.1 Identify where the SFMEA applies
  2.2 Identify the riskiest parts of the software
  2.3 Identify applicable viewpoints
  2.4 Gather documentation and artifacts
  2.5 Identify personnel required
  2.6 Decide selection scheme
  2.7 Set ground rules
  2.8 Review failure severity and likelihood rankings
  2.9 Select template and tools
2.0 Prepare the SFMEA
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Now that you understand the benefits and limitations of the software FMEA and how it’s similar and different than the hardware FMEA, let’s get started with the first step of software FMEA. This step is very important for a successful software FMEA. First, we will identify the scope of the SFMEA so that only the riskiest parts and viewpoints of the software are analyzed. Then we will determine the artifacts and people needed for the particular scope that was previously identified. Different viewpoints require different expertise. Different parts of the software also require different expertise. Once the scope and resources is identified it may/will be necessary to identify a selection scheme for the analysis. For example, you select only 5% of the code for a detailed SFMEA. The last step of the preparation is to tailor the SFMEA to the particular needs of your system and goals. The ground rules are determined to ensure that the analysis doesn’t wonder off from the desired path. The severity and likelihood ratings are defined up front -with respect to your software product - to ensure that they are used appropriately and consistently once the analysis started. The last preparation step is to identify the SFMEA template and tool.

## Slide 14: Identify where the SFMEA applies
The SFMEA can be applied to
  Software components
  Firmware components
  The interfaces to COTS (Commercial Off The Shelf)
  The interfaces to FOSS (Free Open Source Software)
  GFS (Government Furnished Software)
  New and existing code
Identify all software and firmware LRUs in the system
2.1 Identify where the SFMEA applies
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 15: To maximize efficiency and minimize cost
Rank each software LRU based on whether it’s safety critical, safety related or not safety related
Rank each software LRU based on whether it’s mission critical or not
Rank each software LRU based on it’s past or expected failure potential. In other words, identify which software LRUs are likely or known to be more problematic than others.
Sort the software LRUs so that the most safety critical, mission critical and problematic LRUs are at the top of the list
2.2 Identify the riskiest parts of the software
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 16: Identify most relevant viewpoint for a particular software version
2.3 Identify applicable viewpoints
FMEA | When this viewpoint is relevant
Functional | Any new system or any time there is a new or updated set of requirements.
Interface | Anytime there is complex hardware and software interfaces or software to software interfaces.
Detailed | Almost any type of system is applicable. Most useful for mathematically intensive functions.
Maintenance | An older legacy system which is prone to errors whenever changes are made.
Usability | Anytime user misuse can impact the overall system reliability.
Serviceability | Any software that is mass distributed or installed in difficult to service locations.
Vulnerability | The software is at risk from hacking or intentional abuse.
Production | One very serious or costly failure has occurred because of the software. Software is causing the system schedule to slip. Many software failures are being observed at a point in time in which the software should be stable.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: These are the 8 viewpoints and when they are most applicable. Any time you have a brand new software system, the functional viewpoint will be applicable. The only time the functional viewpoint is not applicable is when the code is being changed but the requirements are not changing. An example of this would be if you have product that runs on a particular Operating System and you rewrite the code for the product to work exactly the same but on another Operating System. The code will change but not the software requirements. The interface software FMEA is applicable almost all of the time as it focuses on the interfaces between 2 or more software LRUs or a software LRU and a hardware LRU. The only time an interface software FMEA is less applicable is if the software is very small and it has simple interfaces to very stable hardware. The detailed software FMEA is always applicable. If your system is mathematically intensive this viewpoint may be the most productive at identifying failure modes. However, as we will see later the detailed viewpoint can also be the most time consuming so some sampling is almost always required. The maintenance software FMEA is applicable only when the software is in a maintenance phase of it’s life or if the software is so fragile that any time a change is made to it, a new defect is likely to be introduced. The usability FMEA is most applicable if the user can contribute to a system failure because of the software. The serviceability FMEA applies mostly to software applications that are mass deployed or software applications that are deployed to difficult to reach geography. If the installation package doesn’t work that could mean that many end users, or one difficult to reach end user can’t operate the software. Vulnerability is applicable to most system. It is recommended that your organization seek an expert to help with vulnerability. This presentation provides for failure modes that affect both reliability and vulnerability. However, this presentation does not cover failure modes related to encryption, etc. The production viewpoint is applicable when there are chronic problems with multiple software releases. The goal is to find out what the organization is not developing reliable software as opposed to identifying specific requirements, design, code, install scripts, user manuals, user instructions that can cause the system to fail.

## Slide 17: Failure modes most associated with each viewpoint
Failure mode categories | Description | Functional | Interface | Detailed | Maintenance | Usability | Vulnerability | Serviceability
Faulty functionality | The software provides the incorrect functionality or fails to provide required functionality | X |  | X | X |  |  | 
Faulty timing | The software or parts of it execute too early or too late or the software responds too quickly or too sluggishly | X | X |  |  |  | X | 
Faulty sequence/ order | A particular event is initiated in the incorrect order or not at all. | X | X | X | X |  | X | 
Faulty data | Data is corrupted, incorrect, in the incorrect units, etc. | X | X | X | X |  | X | 
Faulty error detection and/or recovery | Software fails to detect or recover from a failure in the system | X | X | X | X |  | X | 
False alarm | Software detects a failure when there is none | X | X | X | X |  | X | 
Faulty synchronization | The parts of the system aren’t synchronized or communicating. |  | X |  |  |  | X | 
Faulty Logic | There is complex logic and the software executes the incorrect response for a certain set of conditions | X |  | X | X |  | X | 
Faulty Algorithms/ Computations | A formula or set of formulas does not work for all possible inputs | X |  | X | X |  | X | 
2.3 Identify applicable viewpoints
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: At this point you know which viewpoints are applicable, when you can do the SFMEA for that viewpoint and the artifacts you need to collect. In this step we will identify the failure modes usually associated with each viewpoint. The goal of this step is to identify the viewpoints that map to our experience with the most likely failure modes for this type of system or software LRU. On the left column is a list of some common software failure modes. The last 8 columns illustrate which viewpoints each of the failure modes is usually visible. For example if there is a software LRU that is doing GPS, we might want to consider the functional, detailed, maintenance and vulnerability SFMEAs as these pertain to mathematically intensive systems. Another example, you know that in the past this type of software system had problems with synchronization. You might want to consider the interface and vulnerability viewpoints.

## Slide 18: Failure modes most associated with each viewpoint
Failure mode categories | Description | Functional | Interface | Detailed | Maintenance | Usability | Vulnerability | Serviceability
Memory management | The software runs out of memory or runs too slowly |  |  | X | X |  | X | 
User makes mistake | The software fails to prohibit incorrect actions or inputs |  |  |  |  | X |  | 
User can’t recover from mistake | The software fails to recover from incorrect inputs or actions |  |  |  |  | X |  | 
Faulty user instructions | The user manual has the incorrect instructions or is missing instructions needed to operate the software |  |  |  |  | X |  | 
User misuses or abuses | An illegal user is abusing system or a legal user is misusing system |  |  |  |  | X | X | 
Faulty Installation | The software installation package installs or reinstalls the software improperly requiring either a reinstall or a downgrade |  |  |  |  |  | X | X
2.3 Identify applicable viewpoints
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The above shows some more failure modes. Memory management failure modes are typically the most visible when looking at the detailed design or code. Memory failure modes can also result in vulnerability issues. If there have been a considerable number of system failures caused by human being who are attempting to use the software without malice then the usability FMEA may be applicable while the vulnerability FMEA is applicable for malicious users. The next page has even more failure modes…

## Slide 19: Once the viewpoints are selected, these are the artifacts you need to gather
2.4 Gather documentation and artifacts
FMEA | Artifacts that you will analyze
Functional | Software Requirements Specification (SRS) or Systems Requirements Specification (SyRS)
Interface | Interface Design documentation (IDD, IDS)
Detailed | Detailed design (DDD) or code
Maintenance | The code or design that has changed as a result of a corrective action
Usability | Use cases, User’s manuals, User Interface Design documentation
Serviceability | Installation scripts, ReadMe files, Release notes, Service manuals
Vulnerability | See Detailed and Usability
Production | Software schedule, Software process documentation, Software Development Plan (SDP), all development artifacts such as SRS, IDD, IDS, DDD.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Now that you have identified the viewpoints that apply for the particular phase of development that your software is in, you will need to know the artifacts required for the analysis. These artifacts should be requested from the appropriate subject matter experts well in advance to ensure that the software FMEA can be initiated in a timely manner. Either the SRS or SyRS is required for the functional FMEA. The SRS is preferred to the SyRs. The interface viewpoint require an interface design which is usually either a table of interfaces or a diagram. The detailed or vulnerability viewpoint requires either a detailed design or code. Examples of detailed design are state diagrams, timing diagrams, algorithms, user interface diagrams, data flow diagrams, transaction flow diagrams, etc. For the maintenance SFMEA you will need to have access to all of the corrective action reports for the software. If you are performing a usability FMEA you will need to collect the use cases, user’s manuals, and any user design documents. You will need to collect the install scripts, readme files, release notes and services manuals when performing a serviceability FMEA. You will need to collect the software schedules for each individual as well as overall schedules, any software process documents, the software development and all development artifacts shown above for the production SFMEA.

## Slide 20: Depending on the viewpoint these are the people that you will need for the analysis
2.5 Identify personnel required
FMEA | Personnel required for failure mode analysis | Personnel required for Consequences analysis
Functional | Any engineer who understands the requirements for the software | A domain or systems expert who understands the effects of the failure modes
Interface | An engineer who understands the software interfaces | 
Detailed | A software engineer | 
Maintenance | A software engineer | 
Usability | An applications engineer | 
Serviceability | A software engineer | 
Vulnerability | A software engineer | 
Production | Software Management and Software QA and Software Process Group | 
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: Now that you have identified the viewpoints that apply for the particular phase of development that your software is in, you will need to know the artifacts required for the analysis. These artifacts should be requested from the appropriate subject matter experts well in advance to ensure that the software FMEA can be initiated in a timely manner. Either the SRS or SyRS is required for the functional FMEA. The SRS is preferred to the SyRs. The interface viewpoint require an interface design which is usually either a table of interfaces or a diagram. The detailed or vulnerability viewpoint requires either a detailed design or code. Examples of detailed design are state diagrams, timing diagrams, algorithms, user interface diagrams, data flow diagrams, transaction flow diagrams, etc. For the maintenance SFMEA you will need to have access to all of the corrective action reports for the software. If you are performing a usability FMEA you will need to collect the use cases, user’s manuals, and any user design documents. You will need to collect the install scripts, readme files, release notes and services manuals when performing a serviceability FMEA. You will need to collect the software schedules for each individual as well as overall schedules, any software process documents, the software development and all development artifacts shown above for the production SFMEA.

## Slide 21: The scope of the SFMEA may need to be pruned as follows
FMEA viewpoint | Guidelines for pruning | Pruning steps that were taken in section 2.1
Functional | The SRS or SyRS statements that are most critical from either mission or safety standpoint. | The components that perform the most critical functions. The components that have had the most failures in the past. The components that are likely to be the most risky.
Interface | Interfaces relating to critical data or communications. | All interfaces associated with the most critical functions, critical CSCIs or critical hardware.
Detailed | The code that is related to the most critical requirements. Make use of the “80/20” and “50/10” rules of thumb. | The code that has had the most defects in the past. The code that is related to the most critical requirements and CSCIs.
Vulnerability | Identify the weaknesses which are most severe and most likely and look for them in every function | Mitre’s Common Weakness Entry list has ranking. Note that the CWE entries should be sampled and not the code itself. If even one function has a serious weakness then the software can be vulnerable.
Maintenance | All corrective actions in all critical CSCIs | None
Usability | User actions related to critical functions | Safety or mission critical components with a user interface to a human making critical decisions
2.6 Decide selection scheme
Notes: By the time you get to this step in the software FMEA it may be evident that the scope originally identified in section 2.2.1 and 2.2.2 isn’t feasible with the current resources. If that’s the case, this page can be used to prune the scope in such a way as to keep the focus on the high risk areas and failure modes. If the functional software FMEA is selected, the number of requirements identified or the number of failure modes identified for each requirement may need trimming. The interface software FMEA scope can be trimmed similarly by either focusing on some interfaces heavily (with many different failure modes) or focusing on more interfaces with fewer failure modes. The detailed software FMEA almost always requires some type of sampling as analyzing every line of code could be exhaustively expensive. It’s useful to identify the part of the code most associated with the most serious of defects. The vulnerability software FMEA can be trimmed by focusing on the vulnerabilities that rank the highest on Mitre’s Common Weakness Entry AND can be identified via analysis of the detailed design and code. There are many vulnerabilities that cannot be identified via analysis so care should be taken to select those that can. With the maintenance software FMEA, it’s unfortunately not recommended to trim any of the corrective actions. The reason is that even the most trivial corrective action can have huge consequences. The usability software FMEA can be pruned by focusing only on the user actions and interactions that are most associated with mission or safety critical functions.

## Slide 22: Ground rules to consider
Issue | Extent the failure mode is propagated
Human error | Decide whether or not to include human errors in the Functional SFMEAs. The Usability SFMEA focuses on the human error. However, it’s possible to include the human aspect in the Functional SFMEA also.
Chain of interfaces | How many interface chains will we consider in one SFMEA row?
Network availability | Decide whether to assume that any network required for the system is available.
Speed and throughput | Decide whether to assume that the system is performing at maximum, typical or minimum speed and throughput.
2.7 Set the Ground Rules
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The ground rules shown here should be tailored to the particular software LRU or system under analysis. In some cases, human error needs to be part of the analysis while in other cases it may not be. If the interface software FMEA is in scope, you may need to define how many interfaces to analyze at once. You may analyze several in chain at the same time or you may analyze the interface between 2 system components and constrain the analysis to those 2 components. You also need to decide whether to introduce the availability of the network as part of your analysis. Do you assume it’s always available or do you assume that maybe it’s not? The same thing applies to speed and throughput. You need to decide up front whether to assume typical, maximum or minimum speed and throughput and then you need to be consistent in applying that ground rule while analyzing every failure mode.
Review the ground rules
For each item in table on next slide, identify and agree on the ground rules that will be taken when doing this SFMEA
Decide whether to assess the effects, severity and likelihood based on average or worst case. Consistency is important in ranking the likelihood and severity.
Document the ground rules for the SFMEA.
Make sure that all SFMEA participants are aware of the ground rules. During the SFMEA process, the ground rules should displayed in a visible place such as a white board, etc.

## Slide 23: Be familiar with these in advance of doing the SFMEA
2.8 Define failure severity and likelihood ratings
Your organization/project probably has a defined table for assessing likelihood and severity. If not, agree on one.
The below is an example
Agree on concrete definitions for each
  i.e. what has to happen for the failure to be catastrophic?
  This is called a Failure Definition and Scoring Criteria (FDSC)
 | Severity
1 | Catastrophic
2 | Critical
3 | Marginal
4 | Minor
 | Likelihood
1 | Likely
2 | Reasonably Probable
3 | Possible
4 | Remote
5 | Extremely unlikely
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: This looks like an easy activity but it’s often not. Defining the categories for severity and likelihood are not difficult. The difficult part is defining them as per your system. Exactly what is catastrophic – for this system? How does one discern between reasonably probable and possible? The more concrete the definitions are, the easier it will be to perform the analysis. On the other hand, if these definitions are ambiguous it can negatively affect the analysis as well as the results. For example, when the definition of severity is ambiguous it’s not uncommon for all failure modes to be identified as critical or for none of them to be identified as critical.

## Slide 24: Example of an FDSC for a thermostat
Severity | Examples
I | Safety hazard or loss of equipment
II | Persistent loss of temperature control or temperature isn’t controlled within 5% of desired temperature
III | Sporadic loss of temperature control or temperature isn’t controlled within 1 degree but less than 5% of desired temperature
IV | Inconvenience or minor loss of temperature control
2.8 Define failure severity and likelihood ratings
Notes: This is an FDSC for a thermostat. Notice that there are concrete, application specific definitions for each severity level. The definitions should focus on the impact to the system as opposed to the type of defect. For example, a crash does not necessarily have the same severity for every system. For some systems (like a 911 system) a crash may have catastrophic effects while for others (social media) the effect is simply an annoyance.

## Slide 25: Be familiar with mitigation thresholds required for your program or project
2.8 Define failure severity and likelihood ratings
Likely | High | High | Extreme | Extreme
Reasonably Probable | Moderate | High | High | Extreme
Possible | Low | Moderate | High | Extreme
Remote | Low | Low | Moderate | Extreme
Extremely unlikely | Low | Low | Moderate | High
Likelihood/ Severity | Minor | Marginal | Critical | Catastrophic
This is an example of a mitigation risk matrix. The grid is used to determine which failure modes to mitigate and in which order.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: At the end of the software FMEA analysis, the highest ranked failure modes and corrective actions will be reviewed to determine which corrective actions are warranted. Each failure mode/root cause will have an associated Risk Product Number that is simply the severity that you defined multiplied by the likelihood that you defined. As part of the preparation phase, you should determine the shading in the risk matrix. Failure modes associated with cells shaded red are must mitigate, cells shaded orange or mitigate, yellow cells are mitigated when time allows and green aren’t mitigated. The above is an example. The output of this step is to identify the thresholds for mitigation that apply to your product and program. These may already be defined for the hardware FMEA.

## Slide 26: Tips
The same template used for hardware FMEA can be used with some modifications
  The failure mode analysis section will reflect the 8 viewpoints
  You will need to use the template that is associated with the selected viewpoint
  Detectability has a different meaning as discussed later in this presentation
The SFMEA toolkit
  Has all of the viewpoints and failure modes needed for the software FMEA analysis
  Can also be used for hardware FMEA
2.9 Select template and tools
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: At the end of the software FMEA analysis, the highest ranked failure modes and corrective actions will be reviewed to determine which corrective actions are warranted. Each failure mode/root cause will have an associated Risk Product Number that is simply the severity that you defined multiplied by the likelihood that you defined. As part of the preparation phase, you should determine the shading in the risk matrix. Failure modes associated with cells shaded red are must mitigate, cells shaded orange or mitigate, yellow cells are mitigated when time allows and green aren’t mitigated. The above is an example. The output of this step is to identify the thresholds for mitigation that apply to your product and program. These may already be defined for the hardware FMEA.

## Slide 27: Agenda for this section
Analyze failure modes and root causes that apply to each of these viewpoints
  Functional
  Interface
  Detailed design
  Maintenance
  Usability
  Serviceability
  Vulnerability
  Production
3.0 Analyze software failure modes and root causes
There are several hundred possible failure mode/root cause pairs. Just a few will be shown in this presentation for the functional viewpoint. The others are covered in the SFMEA training class and in the SFMEA toolkit.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: There are 8 possible viewpoints for analyzing failure modes. This presentation will cover the first two on the list. The detailed design and maintenance FMEAs are covered in module 2. The usability, serviceability and vulnerability FMEAs are covered in module 3. The production FMEA is covered in module 4. If you have purchased modules 2,3 or 4 you can proceed to those modules once module 1 is completed.

## Slide 28: Functional SFMEA Analysis
Analyze each of the applicable (in scope) software requirements statements for what can go wrong regarding that requirement
Some typical failure modes include but aren’t limited to
  Faulty functionality
  Faulty timing
  Faulty sequencing
  Faulty data
  Faulty error handling
3.1 Functional SFMEA Analysis
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: This analysis can be conducted by software engineers, reliability engineers or systems engineers who are familiar with the requirements of the system. While having software engineering knowledge helps, that’s not required for this viewpoint as long as the analyst is familiar with the system.

## Slide 29: Brainstorm the possible root causes for each applicable failure mode for each SRS statement
3.1 Functional SFMEA Analysis
Failure mode and root cause Section |  |  |  |  | 
SRS number | Related SRS number | SRS Statement | Failure mode | Potential Root cause | Detailed root cause
A reference ID as per the SRS document | List any related requirements by number and text or “none” | Place the statement here | Faulty functionality* | List each root cause (see SFMEA toolkit) | The root cause as it applies to your system
 |  |  | Faulty timing | “” | “”
 |  |  | Faulty sequencing | “” | “”
 |  |  | Faulty data | “” | “”
 |  |  | Faulty error handling* | “” | “”
 |  |  | Others | “” | “”
*Applies to virtually all software requirements
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: These are the steps for performing a functional software FMEA. We will walk through these steps a few steps at a time.

## Slide 30: Example of faulty functionality from history
Satellite is lost at a cost of $186 million.
Engine continues to operate until fuel is consumed
First stage of launch on 10/8/05 is successful. Second stage stops performing when required command to cut off main engine doesn’t occur.
Defect: Flight control system missing a command from the onboard flight control system to the main engine
3.1 Functional SFMEA Analysis
European Space Agency CryoSat-1
It’s unclear why simulator used for testing did not uncover this failure mode. It’s possible that the simulator had the very same fault or that the software testers simply overlooked this fault.
In any case, a “missing” command would certainly be visible during a bottom up review of the requirements, detailed design or code but only if the software engineers are looking at these product documents through the failure space.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 31: Example of timing related failure from history
Therac-25 radiation therapy machine had two modes of operation. First mode is direct electron-beam therapy which delivered low doses of high-energy electrons over short periods of time. Second mode is Megavolt X-ray therapy which delivered X-rays produced by colliding high-energy (25 MeV) electrons into a target.
To protect against over radiation, the second mode required that a beam spreader plate be rotated into place. The system depended on a software only interlock to ensure that the high-power electron beam could only be activated if the beam spreader plate is rotated into place.
3.1 Functional SFMEA Analysis
Patient is over –radiated
System delivers high electron beam with no filter
When operator provided manual inputs at same time as overflow, interlock failed (this is the race condition).
Defect: one byte counter frequently overflowed
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: One of the most famous race conditions in history is the radiation overdoses by the Therac-25 in the 1980s. The radiation overdose occurred because the interlocked failed and the high-power electron beam was activated without the beam spreader plate rotated into place. If the system had had a hardware interlock that would have likely prevented the race condition. The race condition could have also been prevented by writing the code such that it does not allow this important variable to be changed by two sources at the same time. This is also an example of faulty data since the one byte counter was the wrong data size. [THERAC]

## Slide 32: Example of faulty data failure from history - Demonstration of Autonomous Rendezvous Technology (DART)
3.1 Functional SFMEA Analysis
DART (right) used estimates and
measurements to determine its velocity and
position relative to MUBLCOM (left).
Estimates diverged from the measurements and triggered a reset. This imprecision sent DART into an infinite do-loop, resetting itself about once every 3 minutes because of the faulty code that was changed late in testing.
Faulty requirement: The comparison of the velocity input from GPS receiver to software based estimates was specified to have accuracy of ± 2 m/s when it should have been 1 m/s.
Other coding fault: Gain setting code changed right before launch- gave an inappropriately high weighting to the estimates relative to the measurements.
Navigational system could not accurately determine DART’s velocity or position. Excessive (and inaccurate) thruster firings depleted DART’s fuel supply faster than expected.
DART missed the critical waypoint needed to switchover its navigational control completely to the AVGS.
Approximately 11 hours into the 24 hour mission, DART collided with MUBLCOM. Shortly afterwards, its fuel supply ran out, and DART transitioned to its Departure and Retirement sequence.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 33: A very simple example of Faulty Functionality Failure Mode from the Functional (SRS) viewpoint
3.1 Functional SFMEA Analysis
Failure mode and root cause Section |  |  |  |  |  | 
SRS statement number | SRS statement text | Related SRS Statements | Description | Failure mode | Root cause | Detailed root cause
SRS #1 | The software shall display an error message that says “Negative values are not permitted.” whenever a value of <= 0 is entered by the user for the XYZ input field | None | Only values greater than zero are allowed in this input field since XYZ is being used to measure volume. | Faulty functionality | Requirement is missing functionality | SRS statement doesn’t say whether the user is required to acknowledge the message.
 |  |  |  |  |  | The SRS statement doesn’t say what the software is required to do after the message is displayed.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The above example is a simple requirement discussing how the software will handle an erroneous condition. The first root cause for the Faulty Functionality failure mode pertains to what’s missing from this requirement. The analysts review the requirement as well as their understanding of that requirement. They see that 2 things are missing in this requirement. First the requirement doesn’t say whether the user is required to acknowledge the error message. Secondly it doesn’t say what the software should do after this message is displayed. So for the generic root cause “Requirement is missing functionality” there are 2 specific root causes which are added to the FMEA template on 2 different rows. Each of these root causes will then be further analyzed. The next root cause is “Requirement has unwritten assumptions”. The analysts review this root cause and aren’t able to identify any specific root causes so they proceed to the next root cause which is “Conflicting requirements”.

## Slide 34: Continued
3.1 Functional SFMEA Analysis
Failure mode and root cause Section |  |  |  |  |  | 
SRS statement ID | SRS statement text | Related SRS Statements | Description | Failure mode | Root cause | Detailed root cause
SRS #1 | The software shall display an error message that says “Negative values are not permitted.” whenever a value of <= 0 is entered by the user for the XYZ input field | None | Only values greater than zero are allowed in this input field since XYZ is being used to measure volume. | Faulty functionality | Conflicting requirement | One part of the requirement prohibits the value zero while another part allows it.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The next generic root cause for the Faulty Functionality failure mode is “Conflicting requirement”. A conflicting requirement can be across 2 or more requirements or it can be within a requirement. It’s clear when analyzing this requirement for conflicts that it does indeed conflict with itself. The quoted text says that the only negative values are prohibited while the unquoted text includes zero as a prohibited value. It’s not clear which part of the statement is correct. Since the data items is being used to measure presumably it’s the analyst understanding that it can’t be zero. However, the analyst will need to resolve the conflict later in the mitigation phase of the FMEA. The analyst reviews the next root cause which is “Requirement is obsolete”. This requirement is analyzed for obsoletion and it does not appear that obsoletion is relevant to this requirement. So the next root cause is analyzed “Requirement has extra featured.”

## Slide 35: Continued
3.1 Functional SFMEA Analysis
Failure mode and root cause Section |  |  |  |  |  | 
SRS statement ID | SRS statement text | Related SRS Statements | Description | Failure mode | Root cause | Detailed root cause
SRS #1 | The software shall display an error message that says “Negative values are not permitted.” whenever a value of <= 0 is entered by the user for the XYZ input field | None | Only values greater than zero are allowed in this input field since XYZ is being used to measure volume. | Faulty functionality | Requirement has extra features | The message does not have to be displayed if the software doesn’t permit the invalid input.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.
Notes: The last generic root cause for Faulty Functionality is analyzed. At first the analysts do not see how this requirement an have “extra” features. However, eventually that see that it may in fact have extra functionality. The entire requirement is intended to advise the user that they cannot enter negative values. However, the message box may be unnecessary. The user interface can simply not allow invalid inputs. This would eliminate the need for the error message but still leave the requirement that the software not allow the invalid inputs. This would then eliminate the need for the end user to acknowledge the message which is an issue identified earlier in the SFMEA. For now, the analysts don’t attempt to rewrite the requirement, they will save that for later in the mitigation section. For now they identify that the requirement does have an unnecessary feature. At this point the faulty functionality failure mode has been analyzed and the analysts continue to the next failure mode which is faulty timing. Before we proceed to that failure mode it might be useful to see some real life failures that resulted from faulty functionality…

## Slide 36: Summary
The previous example showed a very simple requirement
  Yet there were so many failure modes associated with it
The previous example came from a real SRS for a real product. The specifics have been genericized.
This requirement applied to several hundred inputs.
Imagine just these 4 failure modes repeated hundreds of times for each input.
Imagine what a more complicated requirement or set of related requirements might yield…
The SFMEA toolkit has several real world examples for each viewpoint and each failure mode and numerous root causes
3.1 Functional SFMEA Analysis
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 37: Other failure modes and viewpoints
The SFMEA toolkit has a detailed compilation of failure modes and root causes for the functional viewpoint as well as the other 7 viewpoints.
The analyst can/should also brainstorm additional failure modes/root causes
The analyst should brainstorm all in scope viewpoints and failure modes prior to going to the next step – analyzing consequences.
3.1 Functional SFMEA Analysis
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 38: Agenda
Now that the failure modes and root causes of the applicable viewpoints have been analyzed…
  Identify the subsystem, system effects, effects on operator, etc.
  Identify severity, occurrence and detectability using the ratings in your template
  Preventive measures– These are things that would make this failure mode less or not likely to happen such as hardware interlocks, health monitoring software, etc.
4.0 Analyze Consequences
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 39: Examples of local effects (effect on software itself)
Example of local effects (at the software level)
The wrong function or command is executed
The right function is performed at the wrong time
A commanded function does nothing
Stalls – Take too long
Terminates prematurely
Interruption
Crashes, hangs or freezes
Runs out of memory
Ignores user input
Corrupts data
Loses data
Generates bad data
Generates too much information
Generates stale data
4.1 Identify local, subsystem and system effects
Example of local effects (at the software level)
Behaves erratically
Makes the wrong decisions
Fails to make the correct decisions
Continues processing even when it shouldn’t
Fails to continue processing when it should
Doesn’t restrict user input when it should
Confuses user
Doesn’t work according to the user’s manual
Fails to authenticate end users
Fails to detect security violations or improper authentication
Allows direct access to application memory
Causes end user to become desensitized to real errors
Leaks too much information about how the software works
Allows end user to write data that it shouldn’t
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 40: Examples of subsystem effects
Example of subsystem effects
Loss of subsystem
Loss of required feature
Interruption of subsystem
Degraded subsystem
Incorrect outputs or results from subsystem
Attacker can enter commands instead of data
Attacker can directly access application memory that should be protected
End users ignore errors or relax security because there are too many errors
Error codes aren’t useful to an end user but are useful to an attacker to understand how the software works
Attackers learn about internal state of software from software itself
Attackers can create files in places that typical end users cannot
It’s too difficult for non-malicious users to use
It’s too easy for attackers to get authenticated
4.1 Identify local, subsystem and system effects
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 41: Examples of system level effects (continued)
Example of system level effects
Loss of mission
Loss of equipment
Interruption of service
Degraded service
Injury or safety
Damage to environment
Partial loss of mission
Partial loss of equipment
Loss of product
Loss of security
Loss of revenue
Loss of control over system
Loss of sensitive information
Major annoyance
Inconvenience
Confusion of end user
Loss of private information
4.1 Identify local, subsystem and system effects
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 42: Agenda
Identify corrective actions
Revise Risk Product Number (RPN)
5.0 Identify Mitigation
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 43: Identify Corrective Actions
Corrective action is from the development standpoint and will depend on type of FMEA being performed
  Functional FMEA -corrective action may include changing the specifications
  Interface and detailed FMEAs -corrective action may include changing the design, code to correct the failure mode
  Production FMEA -corrective action may be the execution of a particular practice or avoidance of a particular risk
Examples of corrective actions that don’t apply to software
  Replacing the software unit with an identical failed unit
  Operator repair of the software code on site
Corrective actions must be specific to be effective
  “Improve code review efficiency” is not specific.
  This is specific – “Review the code and make sure that all code that changes global data, locks the data prior to changing it”.
  “Improve testing efficiency” is not specific.
  This is specific – “Input a negative value in field XYZ. Verify that the software doesn’t permit this value to be input and that a user friendly message is displayed”
5.1 Identify Corrective Actions
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 44: Identify Corrective Actions
Corrective actions must be specific to be effective
  “Improve code review efficiency” is not specific.
  This is specific – “Review the code and make sure that all code that changes global data, locks the data prior to changing it”.
  “Improve testing efficiency” is not specific.
  This is specific – “Input a negative value in field XYZ. Verify that the software doesn’t permit this value to be input and that a user friendly message is displayed”
5.1 Identify Corrective Actions
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 45: Revise RPN (Risk Priority Number)
Adjust the RPN based on the assumption that the corrective action is employed
Don’t override the original RPN assessment
5.3 Revise RPN
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 46: Common mistakes made when doing a software FMEA
Personnel are unwilling to view the failure space
Vague failure modes and root causes
Vague corrective actions
Assuming that certain failure modes won’t happen before analyzing them
Analyzing what the code does without regard for what the code is designed to do and is required to do
Focusing at too high or too level a level of abstraction
Neglecting to tailor the list of failure modes to your application type
Neglecting to filter/rank the requirements, design, code, use cases by risk and impact
Analyzing process failure modes instead of product failure modes
6.0 Avoid Common Mistakes
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 47: Common Myths about SFMEAs
Myth #1 – The SFMEA identifies failure modes that we could have found during testing
  There is a very big different between “could have” and “will”.
  Unless the test plan explicitly includes a test for specific negative behaviors – you can’t assume that the resulting failure modes would have been found during testing.
  Test plans rarely cover negative behaviors because they are rarely in the requirements or design documents
  As part of the preparation step the analyst should identify failure modes and root causes that aren’t likely to be immediately visible during testing
Myth #2 – The SFMEA identified a few failure modes. We were expecting many more.
  The goal is to find the failure modes that are difficult to find during testing but can result in one or more very serious events in operation.
  Hence, if the SFMEA finds even 1 very serious failure mode, it can be cost effective.
6.0 Avoid Common Mistakes
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 48: More information
The Software FMEA class covers
  8 viewpoints
  Hundreds of failure modes/root causes
  Real life examples of each
The Software FMEA toolkit has
  A template that starts with preparation and ends with the Critical Items List
  8 worksheets for each viewpoints
  Hundreds of failure modes/root causes prepopulated in the appropriate viewpoint
  Real life examples
  Tips for completing the SFMEA table
Ann Marie Neufelder has been performing SFMEAs for 3 decades on both firmware and software systems for defense, medical, aerospace and energy applications. She can help you perform the SFMEAs.

## Slide 49: References
[1] Delivering Military Software Affordably, Christian Hagen and Jeff Sorenson, Defense AT&L, March-April 2012.
Mil-Std 1629A Procedures for Performing a Failure Mode, Effects and Criticality Analysis, November 24, 1980.
MIL-HDBK-338B, Military Handbook: Electronic Reliability Design Handbook, October 1, 1998.
Society of Automotive Engineers, “SAE ARP 5580 Recommended Failure Modes and Effects Analysis (FMEA) Practices for Non-Automobile Applications”, July, 2001.
NASA, “NASA-STD 8729.1, Planning, Developing and Managing an Effective Reliability and Maintainability (R&M) Program.”, 1998.
NASA, “NASA Guidebook 8719.13B Software Safety”, July, 2004.
W.E Vesely, F.F. Goldberg, N.H Roberts, D.F. Haasl; “Fault Tree Handbook NUREG 0492”, US Nuclear Regulatory Commission, 1981
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.

## Slide 50: References for additional software taxonomies
[ST1] Boris Beizer, “Software Testing Techniques”, Van Nostrand Reinhold, 2nd Edition, June, 1990.
[ST2] Robert V. Binder;”Testing Object Oriented Systems – Models , Patterns and Tools”, Addison-Wesley, 1999.
[ST3] Giri Vijayaraghavan and Cem Kaner, “Bugs in Your Shopping Cart: A Taxonomy”, International Software Quality Week, September, 2002.
[ST4] Giridharan Vilangadu Vijayaraghaven “A Taxonomy of E-commerce Risk and Failures”, A Thesis Submitted to the Department of Computer Science at Florida Institute of Technology, Melbourne, Florida, May 2003.
[ST5] Common Weakness Enumeration, A community developed dictionary of software weakness types. CWE Version 2.6, Edited by: Steven Christey Coley, Ryan P. Glenn, Janis E. Kenderdine, and John M. MazellaProject Lead: Robert A. Martin, © Mitre Corporation, 2014. http://cwe.mitre.org/
[ST6] Peter G Neumann, Donn B Parker; “A summary of computer misuse techniques”, In Proceedings of the 12th National Computer Security Conference, pages 396 – 407. National Institute of Standards, 1989.
[ST7] Ken Erickson, “Asynchronous FPGA risks”, California Institute of Technology, Jet Propulson Laboratory, Pasadena, CA 91109. 2000 MALPD International Conference, September 26-28, 2000.
[ST8] McCall, J.A., W. Randell, and J. Dunham, “Software Reliability, Measurement, and Testing,” Rome Laboratory, RL-TR-92-52, 1992.
[ST9] A. Neufelder, “The Cold Hard Truth About Reliable Software”, Published by Softrel, LLC, 2012.
© SoftRel, LLC 2015 This material may not be reprinted in part or in whole without written permission from Ann Marie Neufelder.